Acquisition of 3D Data  Tomographic Techniques
Acquisition of 3D Data
Tomographic Techniques
Leonid I. Dimitrov and Milos Sramek

---
3D Data
Animation, film: a sequence of images with time as the third coordinate:
[x,y,t]
Volume data: a sequence of images with space as the third coordinate:
[x,y,z]
  - Volume data in general: a set of data samples measured on a grid with certain properties

---
RozptÃ½lenÃ© body
Scattered Points
- Sample position
  - [xi, yj, zk]
- Neighborhood relation:
  - None

---
NeÅ¡truktÃºrovanÃ¡ mrieÅ¾ka
Unstructured Grid
- Sample position
  - [xp, yp, zp]
- Neighborhood relation:
  - Explicit i-k, i-l, …

---
Å truktÃºrovanÃ¡ mrieÅ¾ka
Structured Grid
- Sample position:
  - based on coordinates at [i,j,k]
- Neighborhood relation:
  - Implicit, neighbors from the grid

---
PravouhlÃ¡ mrieÅ¾ka
Rectilinear Grid
- Sample position:
  - Coordinates [i,j,k]
  - Distance between X,Y and Z planes
- Neighborhood relation:
  - Implicit, neighbors from the grid

---
PravidelnÃ¡ mrieÅ¾ka
Regular Grid
- Sample position:
  - Coordinates [i,j,k]
  - Cell size [X ,Y ,Z]
- Neighborhood relation:
  - Implicit, neighbors from the grid

---
KartÃ©zska mrieÅ¾ka
Cartesian Grid
- Sample position:
  - Coordinates [i,j,k]
  - Cell dimension X
- Neighborhood relation:
  - Implicit, neighbors from the grid

---
Cartesian Grid and its Elements
Basic elements of volume data
- Sample:
  - Dimensionless point with value
- Cell:
  - 8 (4) samples in vertices of a cube
- Voxel
  - Homogeneous cube	centered at sample	position
  - Analogue of the 2D pixel
Cell
Sample
Voxel

---
Acquisition of Volumetric Data
Acquisition of Volumetric Data
- 3D imaging techniques
  - Anatomically oriented techniques:
    - Computer tomography-CT, magnetic resonance imaging-MRI, ultrasound imaging-US
  - Physiologically oriented techniques :
    - Positron emission tomography-PET, single-photon emission tomography-SPECT, functional MRI
- Synthetic data
  - Voxelization

---
Computed Tomography
Computed Tomography (CT)
- Also: Computer Aided Tomography (CAT, CAT scan)
- Principle:
  - Measurement of X-ray attenuation along a viewing ray (projections)
  - Production of images: reconstruction from projections

---
CT Tomograph
A CT Tomograph

---
A Sample CT Scan

---
Measurement of Projections
- Standard setup:
  - cca 200 1D projections for one slice
``
Projection
`
Detector
X-ray source
X-ray source

---
Measurement of Projections
- Spatial configuration of source and detector arrays:
  - Aligned rotation
  - 180 deg

---
Fan-Beam Spiral CT
Recent Setup: Fan-Beam Spiral CT

---
CT vs.  Spiral CT
Standard CT vs. Spiral CT
25 Slices in
> 2 Minutes
25 Slices in
25 Seconds

---
Multi-Detector CT Scanner
Multi-Detector CT Scanner
n=256
State of the art:
- Detectors
  - Up to 256
  - Spacing < 1mm
- Scanning speed:
  - 1m/30s
z

---
Reconstruction from Projections
Reconstruction from Projections
- Theory:
  - Johann Radon
- First Tomograph:
  - Godfrey Hounsfield
  - Nobel prize 1979
- Reconstruction methods:
  - Algebraic methods
  - Filtered backprojection
  - Fourier methods
Source: Wikipedia

---
Backprojection
Filtered Backprojection
- For all projections:
  - Projection filtration (high pass filter, derivative)
  - Distribution of the filtered projection to the image in the direction of the projection
All
2
4
8

---
CT Tomogram - an Example
CT Tomogram - an Example

---
Dual Energy CT
- Attenuation coefficient μ depends on the wavelength
- Scanning with two X-ray lamp voltages (energies)
- Advantage: Discrimination of tissue types
Voltage 80kV
Voltage 120kV
Composited

---
Hounsfield Unit
Hounsfield Unit
Normalization of measured values to a -1000 – 3095 (12 bit) scale:
Tissue density is defined by its physical properties

---
CT Data Properties
CT Data Properties
- Axial slices 0.5-10 mm thick
- Images 256 x 256 to 512 x 512 pixels of 0.5-2 mm side
- Up to 2000 images per study
- High spatial resolution
- Consistent values (HU scale)
- X-ray irradiation

---
Application Fields
Application Fields
- Contrast determined by tissue density:
  - Bone imaging
  - Calcifications
  - Inflammations
  - Hematomata
  - Tumors
- Imaging with contract agent:
  - Blood vessels, vessel lumen

---
Overview
- Computer tomography (CT, CAT)
- Magnetic resonance imaging (MRI)
- fMRI
- SPECT
- PET

---
Magnetic Resonance Imaging
Magnetic Resonance Imaging
- Nuclear Magnetic Resonance (NMR) Magnetic Resonance Imaging (MRI)
- Physical principle
  - Interaction of atom nuclei with an external magnetic field (resonance)
  - Requirement: Nonzero magnetic or spin moment of the atoms
- Resonance
  - Energy transfer between coupled systems with equal characteristic frequency
  - Example: The Tacoma Narrows bridge

---
PowerPoint Presentation
Common NMR Active Nuclei
Isotope	Spin	%	
  I	abundance	MHz/T
1H	1/2	99.985	42.575
2H	1	0.015	6.53
13C	1/2	1.108	10.71
14N	1	99.63	3.078
15N	1/2	0.37	4.32
17O	5/2	0.037	5.77
19F	1/2	100	40.08
23Na	3/2	100	11.27
31P	1/2	100	17.25

---
Spins in a Magnetic Field
Spins in a Magnetic Field
M 0
M=0
spin
B0 = 0
B0  0
Atoms with uncompensated
magnetic moment
M - Magnetization (material)
B - Induction (field strength)

---
Energy in a Magnetic Field (Zeeman Splitting, Spin ½)
Spin in a Magnetic Field
(Zeeman Splitting, Spin ½)
B0
B0
mI = +½
mI = ½
1.5T, T=310K,
P+1/2= 0. 5000049
P-1/2= 0.4999951
 Total magnetization M is parallel to B0

---
Spin in a Magnetic Field
Spin in a Magnetic Field
-  rotates around B0
- B0  external mg. field
-  spin moment
  - rotates around B0 with Larmor frequency
        fL =  B0
-  gyromagnetic ratio (= 42.58 MHz/T for H)
- Total magnetization:
        M =  μi
B0

fL

---
MRI System Block Diagram
MRI System Block Diagram
B0
gradient field
coil
magnet
radiofrequency coil
RF
amp
spectrometer
X
amp
Y
amp
Z
amp
z
y
x

---
Magnetic Field of a Tomograph
Magnetic Field of a Tomograph
- Induced magnetic field:
  - Electromagnets
  - Superconductive electromagnets
  - Permanent magnets
- Field strength
  - 0.5T – 3T, (up to 15T for research)
- Earth’s mg. field : 0.3 - 0.7 x 10-5T

---
Principle of MR Imaging
Principle of MR Imaging
Absorption and emission of energy by spins in external magnetic field (resonance)
Equilibrium: M || B0: Uncoupled μi vectors:
        Mz = M, Mx y = 0
- Absorption of RF (radiofrequency) energy: excitation(f = fL)
  - Transition to higher energy state: Mz< M owing to vector μi flipping
  - Coupled rotation of μi: Mx y  0
  - Measurable signal due to Mx y
- Relaxation: return to equilibrium

---
Spin Excitation
Spin Excitation and Relaxation
B0
B0
M
M
M
Excitation by RF impulse
Relaxation: Return back to equilibrium.
Equilibrium

---
Relaxation
Relaxation
Return to equilibrium state after the RF pulse
- Longitudinal relaxation (spin-lattice, rate T1):
  - Energy transfer to surrounding tissue
  - Flipping of μi vectors to original orientation
- Transversal relaxation (spin-spin, rate T2):
  - Decoupling of i due to field inhomogeneities and spin-spin interaction
- Measured signal during the relaxation
  - FID (Free Induction Decay)
t

---
Tissue Contrast
Tissue Contrast
Both T1 and T2 depend on the chemical environment of the nucleus (spin).
M
Contrast
Tissue A
Tissue B
t
TE: time of measurement, echo time

---
Tissue Contrast
Effect of TE (echo time, time of measurement)

---
Tissue Contrast
Effect of echo time

---
Tissue Contrast
Effect of echo time

---
Tissue Contrast
Effect of echo time

---
What Effect Has T1 on Images?
MRI Measurement
RF pulse (excitation)
Mz
TR
TE
- TR: repetition time
- TE: excitation time, moment of measurement

---
T1, T2 and PD Images
T1, T2 and PD Images
- Different combinations of TR and TE yield different tissue contrast:
  - Long TR (>1500ms)
   Short TE (<25ms) PD (Proton Density)
        weighting (no T1 and T2 influence)
Long TE (>50ms) T2 weighting
  - Short TR (<500ms)
   Short TE (<25ms) T1 weighting

---
T1, T2 and PD Images
T1
T2
PD

---
MRI Overview
- Physical Background
- Excitation and Relaxation
  - T1 Relaxation
  - Spin Echo and T2 Relaxation
- Coding of Spatial Information
  - Slice Selection
  - Read-out Gradient
  - Phase Encoding

---
Spatial Localization in MRI
Spatial Localization in MRI
- Position encoding by means of gradient fields:
  - Z-gradient: slice selection by changing the Larmor frequency fL = (B0 + Gz.z)
  - Z-gradient is applied
   during the RF pulse

---
Slice Selection
Slice Selection

RF pulse
frequency
0
time
+
Gz

---
Read-Out Gradient
X, Y encoding
- Similar tricks to encode the x and y coordinates
  - Only one row measured in one excitation
Bz
x

---
Examination Time
Examination Time
- Image is measured row-by-row
- Basic formula: T = TR x R x N [ms]
  - TR : repetition time
  - R : number of image rows
  - N : Number of accumulations (noise)
Example: T = 2000 x 256 x 2 = 17 min
- Speed-up:
  - Low excitation angle  shorter TR (low energy pulse)
  - Multislice techniques

---
Multislice Technique
Multislice Technique
Interleaved excitation of several (64) slices:
TR
Slice 1
TR
Slice 2
TR
Slice 3
TR
Slice 4
TR
Slice 5

---
Imaging in Other Planes
Imaging in Other Planes
- B0 always remains in the same direction
- Choice of imaging plane depends on the order of gradients’ application
- Oblique planes: simultaneous application of 2 gradients
Sagittal
Coronal
Transversal

---
Scanning Protocols
Scanning Protocols
- Protocol: a sequence of pulses, gradients and signal measurements
- Protocols influence image properties and examination time
- Patented and sold
    by scanner vendors
FISP, FAST, FLASH,
   STAGE, STERF
t
s
B1
FID
Echo
Gx
Gy

---
Properties of MR data
Properties of MR data
- Measurement in arbitrary planes
- Typically 256x256 (512x512)pixels
- No absolute scale for measured values
- Significant level of noise
- Good soft tissue contrast
- Spatial inhomogeneities - bias
- No irradiation

---
MR data - examples
MR data - examples
T1
PD
MRA

---
Overview
- Computer tomography (CT, CAT)
- Magnetic resonance imaging (MRI)
- fMRI
- SPECT
- PET

---
Functional MRI (fMRI)
Functional MRI (fMRI)
- Assumption: Active brain area needs more oxygen than inactive ones
- fMRI: statistical detection of areas with oxygenated blood flow
- Visualization by merging with regular MRI data

---
fMRI Example
fMRI Example
Right hand finger tapping at 2Hz

---
Overview
- Computer tomography (CT, CAT)
- Magnetic resonance imaging (MRI)
- fMRI
- Emission Tomography
  - SPECT
  - PET

---
Emission Tomography
Emission Tomography
ECT - Emission computed tomography
- Introduction of a radioactive agent into the patient’s body (tagged substance)
- Agent’s distribution based on metabolism
- Measurement of its spatial distribution

---
SPECT
SPECT
Single-Photon Emission Computed Tomography
- Isotopes emitting photons (Tc-99, I-125, I-131)
- Uniform distribution of photons in all directions
- Scanner - a set of detectors with collimators - gamma camera

---
Gamma Camera
Gamma Camera

---
Gamma Camera

---
SPECT data
SPECT data

---
CT-SPECT
- Combination of different techniques – supplementary information
- Registration is required
  - Not a trivial task
  - Solution: CT-SPECT scanner

---
CT-SPECT (2)
CT
SPECT

---
CT-SPECT (3)

---
PET
PET
Positron Emission Tomography
- + decay (positrons)
- Annihilation - a pair of photons with energy 511keV in opposite directions
- Detection without collimators: registration of concurrent events in detector pairs
- 3D measurement, statistical reconstruction (MRF)

---
PET Scanner

---
PET data
PET data

---
PET Scanners
Research prototype
at BNL (1961)
PET today

---
Combined PET/CT, MRI
PET / MRI
PET / CT

---
Other Imaging Modalities
- Ultrasound
- Confocal microscopy
- Electrical impedance tomography
- Optical coherence tomography
- ...

---
Overview
- Computed tomography (CT / CAT)
- Magnetic resonance imaging (MRI)
- fMRI
- SPECT
- PET
- Synthetic Data

---
VoxelizÃ¡cia geometrickÃ½ch objektov
Voxelization of Geometric Objects
- Preparation of “synthetic” data
- A CG alternative
- Simultaneous visualization of geometric and volume data
- Modeling of volumetric properties (e.g. weathering)

---
Surface  and  Volume  Graphics
Surface and Volume Graphics
- Surface Graphics: Geometric rendering of Continuous Spatial Models (CSM)
- Volume Graphics: Voxelization of a CSM, manipulation, and volume rendering of a Volumetric Model (VM)
CSM
Picture
VM
Geometric
rendering
Volume
rendering
Voxelization

---
Voxelization
Voxelization
- A process of approximating a continuous geometric primitive in the 3D discrete space
- The result of this process
Binary voxelization
Voxels & occupancy O {0,1} ({Background, Foreground}, {Black, White})
Non-binary voxelization (fuzzy, filtered)
Grid points & densities D  R

---
Normal  Estimation in  Binary  Data
Binary vs. nonbinary
Binary
Non binary

---
Non-binary Voxelization
Suppression of aliasing in
3D model by smooth
density transition in the
surface area
- Truncated distance field
- Surface reconstruction
  - Interpolation and thresholding
Analytic object
Voxelized object

---
Voxelization  of  Solids  &  Surfaces
Solids in Truncated Distance Fields
Object
surface
1
0.5
Density transition
area
2
 Surface density profile

---
Voxelization by Direct Distance
Computation
Simple primitives (sphere, torus, polygon)
Example: Voxelization of a triangle:
Bounding box
Voxel-by-voxel update of distances to plane, edges & vertices
Density according to the minimal distance

---
Voxelized  Polygonal  Model
Voxelized Polygonal Model

---
Voxelization of Parametric
Surfaces
P(u,v) = [ x(u,v), y(u,v), z(u,v) ]
[ u, v ]  (u0, u1) x (v0, v1)
1. Splatting – adding small voxelized balls
2. Approx uniform sampling
by binary domain subdivision
Domain
subdivis.
Surface
subdivision

---
Voxelization of Parametric
Surfaces

---
Voxelization of Implicit Solids
{[x,y,z]: f(x, y, z) < 0}
Distance estimation by linear approximation:
d(x,y,z)={f(x,y,z)} over { ldline f'(x,y,z) rdline }

---
Voxelization  of  Implicit  Solids
Voxelization of Implicit Solids
(2x2 + y2 + z2)3 - (x2/10 + y2)z3
(y2 + s)(x2 + z2) - s

---
CSG Operations

---
CSG  Operations
CSG Operations
a * (b + c)
a
b
c